	[image: image1.png]CESKA
o o:chooni
INSPEKCE




	PRESS RELEASE


The KN95 FFP2 respirator does not meet the requirements of the FFP2 type respirator 

(Prague, August 3, 2023) The Czech Trade Inspection Authority found that a respirator on the domestic market does not meet the requirements for FFP2 respirators.

The Czech Trade Inspection Authority found that a model of personal respiratory protective equipment is offered on the Czech market for consumers that does not meet the requirements of Regulation (EU) 2016/425 of the European Parliament and of the Council on personal protective equipment and repealing Council Directive 89/686/EEC. The product was offered under the name KN95 FFP2 Protective Mask at the time of the inspection. It is a white filtering half mask with an embedded wire for shaping around the nose. The product was sold individually in transparent plastic packaging without printing and has no markings other than the embossed KN95 marking. Instructions for use were not included.

Laboratory tests were carried out in an accredited laboratory to verify the total permeation, permeation through the filter material and breathing resistance according to the requirements of EN 149:2001+A1:2009, Articles 7.9.1, 7.9.2 and 7.16. The results of the tests carried out show that the product did not meet the requirements for total permeation and permeation through the filter material. The product does not comply with the requirements of EP and Council Regulation 2016/425 on personal protective equipment. Due to the fact that the filtering half mask has a higher total penetration and penetration through the filter material, it does not meet the requirements for FFP2 type respirators and cannot effectively protect the user from solid particles or liquid aerosols. The product also fails to meet the requirements for labelling, instructions and consumer information. Furthermore, the EU declaration of conformity confirming that the product has followed the conformity assessment procedure in accordance with the requirements of the above mentioned Regulation was not provided.     

The Czech Trade Inspection Authority has imposed a protective measure on the distributor of the product, which is a ban on the supply of the respirator in question to the market and its use for own consumption, including its withdrawal from the market and circulation. Information on the identified product deficiencies was reported to the MIT contact point and published in the Safety Gate information system.  
Product information: 

Product namee:



KN95 

Trademark:



not displayed
Model Type / Number:


not displayed
Production serie number:

not displayed 

Identification of importer:

Import World s.r.o., Lannova 2061/8, 110 00 Praha, 

                                                                        IČO 02979161

Identification of manufacturer:: 

Zhejiang Steriger Sports Safety Technology Co., Ltd., Houzhai 

                                                                        Industrial Zone Yiwu Zhejiang

	[image: image2.jpg]



	[image: image3.jpg]





	  Picture of the product packaging
	Picture of the product



Contact:
Spokesperson of the CTIA
                                                            Telefon:
+420 296 366 233  
František Kotrba
                                                            Mobil: 
+420 731 553 732

                                                            E-mail:
mluvci@coi.cz

